Table of Contents

Quick start guide	2

Instrument start-up	2

Sample preparation	3

Creating new experiment from template	3

Running your samples	4

Analyzing fcs files	4

Shut Down & Clean Up	4

Maintenance	5

Troubleshooting	5

# Quick start guide

# Instrument start-up

- Turn on power to the instrument

- Fill out the sign-in sheet on the google drive

- Open NovoExpress

- Prepare the QC Particle Sample (not needed every run, but before big experiments or if not done within the same day)

  - 0.2 um sterilize TE (and MilliQ?)

  - Mix 0.8 ml of TE with 0.2 ml of MilliQ water in a flow cytometer tube

  - Shake or vortex the QC particles and add 2 drops to the TE+MilliQ mix

  - Vortex mixture

- Go to Instrument>QC test

  - Make sure QC particle tube is in well A1

  - Make sure QC particle Lot ID is entered in correctly

  - Check for system test

  - Prepare sample “Yes”

  - Run

  - QC results

    - If passed, move on

    - If it did not pass, run ‘expansive cleaning sequence’ – steps are in a file on the desktop labelled ‘recommended cleaning sequences’. Ensure cables are tightly secured. If it fails again, ask for technical support.

# Sample preparation

- In the dark, make a 50x working stock of SYBR Green I: Dilute 5 ul of 10,000x stock into 995 ul of 1x TE (0.2um filtered)

- Fix samples with glutaraldehyde at a final concentration of 0.25%

- Stain samples with a final concentration of 0.5x SYBR Green I

- Make sure to include a stained TE (0.2um filtered) and/or seawater (0.2um filtered) blank!

- Incubate samples in the dark at room temperature for 15 – 30 min

- Prior to running samples, serially dilute in either 1x TE (0.2um filtered) or seawater (0.2um filtered)

  - Samples with virus and bacteria get diluted in TE

  - Phytoplankton (and environmental?) samples get diluted in seawater

# Creating new experiment from template

- Create a New Experiment from Template

  - For live phytoplankton, click the ‘Live Phytoplankton Template’

  - For SYBR Green stained samples with viruses and bacteria, click the ‘Virus Bacteria SYBR Green Template’

- Double check that your flow cytometry parameters are correct.

  - Either view a previous successful experiment or file_XXXX for gain parameters, flow speeds, and thresholding

  - If you change any parameters for a single sample, make sure in the Experiment Manager on the right side toolbar that you drag the ‘Cytometer Settings’ for that sample to the Specimen1 to update all samples with these settings, and double check.

- Navigate to the Experiment Manager on the right toolbar, and click on ‘Work List’

  - Here you can rename your samples

  - To import a pre-made spreadsheet, do XXXX

  - Delete samples in slots you will not be needing

  - Click ‘Apply Modification’

- Create plots for your type of data (note: we may edit the template so that all the plots one might need are there). If you make a plot for a single sample, drag the ‘Analysis’ icon in the Experiment Manager onto Specimen1 to update all samples with the plots.

  - Make sure x and y axis are set to log and show the full range

  - Right click on x or y axis labels to change the channel displayed

  - Make sure to open at least one plot that has flow? On the x-axis and counts on the y-axis. This will tell you if there were any bubbles or things getting stuck in the SIP

# Running your samples

- Arrange your prepared samples into the appropriate tube rack

- It is good practice to label your samples and create the appropriate plots for your experiment

  - B581 vs FSC for XXX

  - B680something vs FSC for XXX

  - Counts vs time to catch any bubbles in the system - it should be a steady stream of events. Bubbles will compromise the data?

- It is good practice to create a plot for [] x time so you can catch any bubbles in the system as these will appear as blanks in the plotted events.

- Refer to the Shut down & clean section when finished

# Analyzing fcs files

- Choice of software

  - Free software in browser https://floreada.io/

  - Using the NovoExpress software (only on flow cytometry computer or lab laptop

  - FlowJo (we have a license?)

- Something about gates and controls

- Something else probably

# Shut Down & Clean Up

Click on ‘Instrument’, ‘Shut Down’ and follow either IF statement below -

IF you analyzed ocean water samples and/or the instrument was left idle for >1h as ocean water might have dried off at the SIP (Sample Injection Probe) and salt water crystals might have formed:

- Select for ‘manual cleaning of the SIP’. The software will instruct you how to proceed.

- Dampen a kimwipe with NovoClean and gently wipe the surface of the SIP, the software will prompt you to click “OK” when you are done.

- Fill a sample tube with 4 mL NovoRinse (MilliQ water), which will be used to submerge the SIP. We have solutions for SIP rinsing (0.22 µm-filtered DI water or Novo Rinse).

  - The SIP will remain submerged in the SIP rinse fluid tube and it will retract the next time you start the instrument.

  - If you need to reclaim the tube rack and/or just want to pull it out – you can open the top cover on the Opteon and just manually lift the SIP assembly straight up to pull it back. Once the instrument is powered off, it releases control of the SIP and it should slide back up fairly easily.

IF you analyzed ocean water samples and the instrument was left idle for <1h reference 2.

- Do not select for ‘manual cleaning of the SIP’

- Run 2-3 DI sample tubes (at the same or slightly above water volume than the actual/ocean water samples; likely 4-5 ml will work) at the end of the analysis to clean up the instrument and rinse the SIP.

# Maintenance

- Please see Agilent FCM: Maintenance

# Troubleshooting