Quixstand TFF System Protocol

Terminology:

- Cross-flow/Tangential flow = sweeping action created by fluid flow across the membrane

- Feed stream/Process solution = bulk solution to be processed by the cartridge

- Retentate/Concentrate/Reject = solution containing species retained by the membrane

- Permeate/Filtrate/Ultrafiltrate = solution containing solvent and solutes passing through the membrane

Shear rates:

According to the manual, shear rates of 2000 – 4000 sec-1 are appropriate for viruses, less probably won’t hurt

Shear rates of 4000 – 6000 sec-1, and 8000 – 16000 sec-1 are used after sample concentration for cleaning

- For the 4M or 4X2M housing [the MA on our carts has an A for autoclavable] , inner diameter (ID) 0.5 mm

  - Shear 2000 sec-1 0.3 L/min

  - Shear 4000 sec-1 0.6 L/min

  - Shear 8000 sec-1 1.2 L/min

  - Shear 16000 sec-1 2.4 L/min

To measure shear rates, measure the flow rate of both the permeate and retentate using a graduated cylinder. Add the values together to get the feed flow rate

Opening a new cartridge:

Cartridge contains glycerol that needs to be flushed out. First stage is an EtOH rinse. According to ‘Hollow fiber cartridge preparation instructions’ that came with the UFP-300-C-4X2MA cart, alcohol rinse used is 25% EtOH

Set a water bath at 50C with a few 1 L flasks filled with UltraPure MilliQ Water

Alcohol rinse:

- Connect cartridge to the system. Permeate is preferred to come out of the lower port

- Fully open the retentate and permeate ports

- Connect the retentate and permeate lines to the feed reservoir containing 25% EtOH

- Start the pump on slow (I did 30 RPM)

- Begin closing the retentate and permeate lines so that the transmembrane pressure (TMP) is ~5 psig. Take note that our pressure gauge 0 is stuck around 14 psig. Yellow tape indicates 0, 5, and 10 psig

- Recirculate the 25% EtOH for 10 mins

- Turn off the pump and allow the cartridge to soak in place for 30 min

- Transfer the 25% EtOH from the reservoir to a new bottle – original manual says this can be saved and re-used several times before discarding

- Fully open the retentate and permeate lines to drain them to the appropriate waste container (sink is fine for dilute EtOH)

Water rinse:

- Fill the feed reservoir with ultrapure MilliQ water that is room temperature or warm (up to 50C). Do NOT use cold water

- Retentate and permeate lines should be fully open and should drain directly into the sink

- Start the pump on slow (I did 25 RPM)

- Adjust the TMP to 5 psig by closing the retentate and permeate lines

- Keep adjusting until TMP is 5psig, and retentate flow is 1/10th of the permeate flow so that most of the water is getting pushed out of the cartridge. I used a graduated cylinder to first measure the flow over the permeate line, then the retentate line separately.

- Continue to rinse for 90 mins. This will require topping off your feed reservoir many times, make sure you are near the TFF unit and have at least 2 x 1 L of ultrapure MilliQ water in the 50C water bath ready to be used – once you use 1 L, fill it up with fresh MilliQ and place into the water bath to be ready for the next top-off. If the water is not cold this should be fine

Clean water flux evaluation:

- These should be able to be measured at easily repeatable conditions. I set the small green Watson marlow pump 323 at 50 RPM

- I used fresh MilliQ water direct from the system

- I cracked the retentate line so that there was a small drip flow and opened up the permeate line until I was at a TMP of 5 psig

- For the UFP-300-C-4MA, I got a clean water flux of 158 ml/min

  - This is a 650 sq cm surface area, or 0.065 sq m

  - Flux (liters per surface area per h[lmh])= permeate flow(ml/min)/cartridge area (m2) * 0.06

  - Flux of UFP-300-C-4MA = 145 lmh

  - It is really stupid that the manual tells you to make this conversion, then on page 32 they show a table in L/min. This table also is hard to understand what the numbers on the top mean.

- Remove the cartridge from the quixstand either prepare to concentrate sample, or store cartridge.

Storage: Store cartridge fully submerged (if possible) in 0.1 N NaOH. If too tall, make sure to parafilm any open ports so that the cartridge does not dry out